## Slide 1
Figure 4: First row: ground truth data. Second row: reconstruction error using filtered backprojection. Third row: reconstruction error with proposed iterative method and polychromatic forward model.
References
[1] N. Bevins et al.: “Beam hardening in x-ray differential phase contrast computed tomography,” in Medical Imaging, March 2011.
[2] G. Pelzer et al.: “A beam hardening and dispersion correction for x-ray dark-field radiography,” Medical Physics, vol. 43, May 2016.
[3] A. Ritter et al.: “Simultaneous maximum-likelihood reconstruction for x-ray grating based phasecontrast tomography avoiding intermediate phase retrieval,” arXiv:1307.7912, Jul. 2013.
Conclusions
Reconstruction method for Polychromatic Talbot-Lau imaging
Experiments with synthetic phantom show strong artifact reduction
Efficient optimization method for phase-image required
Further experiments with real data in progress
Polychromatic Maximum Likelihood Reconstruction for Talbot-Lau X-ray Tomography
Florian Schiffers1,2 , Sebastian Käppler1, Shiyang Hu1, Jens Rieger2, Georg Pelzer2, Andreas Wolf2,
Andreas Maier1, Thilo Michel2, Christian Riess1, Gisela Anton2
1 Pattern Recognition Lab, Department of Computer Science, Friedrich-Alexander University Erlangen-Nuremberg, Erlangen, Germany 2 Erlangen Centre for Astroparticle Physics, Department of Physics, Friedrich-Alexander University Erlangen-Nuremberg, Erlangen, Germany
Materials and Methods
Goal: Simultaneous reconstruction [3] of attenuation, phase and scatter coefficients inherently removing polychromatic artifacts
Estimate phase-stepping data (Fig. 1) seen by pixel at stepping position
Reference images , , and obtained using wave propagation simulation (Fig. 2)
Attenuation , differential phase , and dark-field at energy given by
System matrix elements and calculated by integration over footprint of Kaiser-Bessel function (Fig. 3)
), ), and extrapolated using a physical model
Minimization of negative log-likelihood
Introduction
Polychromatic artifacts in Talbot-Lau tomography due to energy dependence of
Material: Attenuation coefficient , refractive decrement [1], and scatter coefficient [2]
Talbot-Lau interferometer: Spectrum , spectral visibility and energy dependent reference phase
Clinical application requires beam hardening correction
Results and Discussion
Synthetic CT data affected by polychromatic effects generated from well-defined phantom using proposed forward model (Fig. 4)
Reconstruction removes polychromatic artifacts (Fig. 5)
Slow convergence of phase-image
Contact
 Florian.Schiffers@fau.de, Sebastian.Kaeppler@fau.de
Figure 5: Two cross-sections through the reconstructed data as visualized in Fig. 4. Black: Phantom ground truth, Blue: Filtered Backprojection, Red (dashed): Proposed reconstruction method. For all images, the iteratively reconstruction and ground truth are in high agreement.
FBP
Ground Truth
Iterative Rec.
Attenuation
Refractive decrement
Scatter
H2O
PTFE
PMMA
0
20
40
60
20
40
60
5
10
15
20
-10
0
10
0
2
0
0
2
0
4
-0.2
-0.1
0
0.1
0.2
-1
-0.5
0
0.5
-2
-4
1
-0.4
-0.2
0
0.2
-2
-4
0.4
4
(a) Attenuation
(b) Refractive decrement
(c) Scatter
Figure 1: Superposition of mutually independent forward projections for phase-stepping data (here shown for phase step 0) at different energies result in the expected phase-stepping data in a polychromatic setup.
15 keV
30 keV
45 keV
60 keV
Polychromatic
Figure 2: Spectral visibility and X-ray spectrum for the interferometer used in the evaluation.
Figure 3: (Differential) footprint of Kaiser-Bessel function.